## Slide 1
Recent InGrid Studies and DESY Test-Beam
Results with Octopuce
Maxim Titov, CEA Saclay, Irfu/SPP, France
Ingrid
Workshop LAL-Ukraine on instrumentation LAL Orsay, France, May 22-24, 2013
University of Bonn: Y. Bilevych, C. Brezina, K. Desch, J.
Kaminski, T. Krautscheid, C. Krieger, M. Lupberger
NIKHEF: Jan Timmermann
CEA Saclay: David Attie, Xavier Coppolani, Andrii Chaus, Paul Colas, Maxim Titov
+ New Teams interested in Si Pixel –TPC R&D :
DESY Hamburg (T. Behnke et al.), LAL Orsay (S. Barsuk et al.), Kyiv University (O. Bezshyyko et al.)

## Slide 2
+
-
+
-
Medipix2/
TimePix chip
X-ray source
Ionizing
particle
Flip-chip
bump bonding
connections
Semiconductor
sensor
Gas volume
Amplification System (MPGD)
Solid detector
Gas
detector
Pixel Readout of Micro-Pattern Gas Detectors
Gas Detector Readout by multi-pixel CMOS array (used as charge collecting anode)
CMOS readout concept  MEDIPIX2 / TIMEPIX Chips
Medipix2/Timepix
ILC-TPC: Ultimate Gas-Silicon Detector Integration
First studies of MPGDs with Medipix2:
M. Campbell et al, NIMA 540, 295 (2005).
M. Chefdeville et al, NIMA556(2006) 490
(NIKHEF-Saclay-CERN- Univ. Twente/MESA+)

## Slide 3
Each pixel can be set to:
● TOT ≈ integrated charge
● TIME = Time between hit and shutter end
“InGrid” Concept:
Protection Layer (few mm)
50 - 80 mm
(e.g. Si3N4)
(e.g. SU8)
Bump bond pads for Si-pixel
Detectors - Timepix or
Medipix2 (256 × 256 pixels
of size 55 × 55 μm2) serve
as charge collection pads.
3D Gaseous Pixel Detector  2D (CMOS pixel chip readout) x 1D (drift time)
Integrated Electronics: Pixel Readout of Micro-Pattern Gas Detectors
Through POST-PROCESSING INTEGRATE
MICROMEGAS directly on top of CMOS chip

## Slide 4
"InGrid" Detector: Single Electron Response and Discharges
Fe55
source
Observe electrons (~220) from an X-ray (5.9 keV) conversion one by one and count them in micro-TPC (6 cm drift)  Study single electron response
1.5 cm
P. Colas, RD51 Collab. Meet.,
Jun.16-17, 2009, WG2 Meeting
Provoke discharges by introducing small amount of Thorium in the Ar gas - Thorium decays to Radon 222 which emits 2 alphas of 6.3 & 6.8 MeV
Chips survive thousands of discharges
M. Fransen, RD51 Collab. Meet.,
Oct.13-15, 2008, WG2 Meeting
Saclay
NIKHEF

## Slide 5
Two micro-TPC boxes have been built
Drift distance in micro-TPC (~ 10 cm)
is large enough to allow study of single
electron response from Fe55 source
Saclay Micro-TPC with Timepix
TIMEPIX
Micro-TPC
Fe55

## Slide 6
Studies of IZM “InGrid”s in the Saclay micro-TPC
Studies of acceptance tests of newly produced “InGrids”
See O. Bezshyyko talk  simulation
of “InGrid” performance
in micro-TPC and
and comparison with
with experimental measurements
Guard ring issue
in micro-TPC
Some noise and/or
discharges
Some local grid
issues

## Slide 7
“InGrid’ Technology and “Driving” Developments
2011: Major Step Forward 
InGrid Production on a wafer level (107 chips)
2005: Single “InGrid” Production
2009: “InGrid” Production on a
3 x 3 Timepix Matrix

## Slide 10
Saclay “Octopuce” (2009)
HV connections
“Grid irragularities”
Non homogeneous area on the mesh of different chips
chip 1
chip 4
chip 5
chip 8

## Slide 11
11
Chip 1
Saclay Octopuce & Fe55: Homogeneous Irradiation
(pixels)
(sum of
pixels TOT)
Fe55 Studies (26/07/2012):
He/Iso 80/20
Vmesh=390 V
Vdrift =3000 V
Integrated picture for all 8 chips:
(discontinuity = distance
between pixels in the cluster)
Chip 4
“Hot spots”
Chip 5
Chip 8
A.Chaus, D. Attie

## Slide 12
Saclay Octopuce: Uniformity of the Response (I)
THRESHOLD = 1000 COUNTS
THRESHOLD = 2000 COUNTS
THRESHOLD = 4000 COUNTS
Fe55 Studies (26/07/2012):
He/Iso 80/20
Vmesh=390 V
Vdrift =3000 V
Non-sensitive (~ 1.5 mm)
areas between chips
Chip 1
Chip 4
Chip 5
Chip 8
A.Chaus, D. Attie

## Slide 13
Saclay Octopuce: Uniformity of the Response (II)
Non-sensitive (~ 1.5 mm)
areas between chips:
THRESHOLD = 6000 COUNTS
THRESHOLD = 8000 COUNTS
THRESHOLD = 10000 COUNTS
More studies required to understand:
if “dead” areas between chips decreases
with increased drift field;
“dead” area on the outer edge as a
function of the guard voltage
A.Chaus, D. Attie

## Slide 14
Expected number of primary electrons in He/Iso (80/20) ~ 165
Saclay Octopuce: Cluster Size Distribution
Single electron sensitivity is very high for all (but 1 and 4) chips
A.Chaus, D. Attie

## Slide 15
Chip number | Threshold level
1 | 337
2 | 335
3 | 340
4 | 335
5 | 330
6 | 327
7 | 330
8 | 320
Saclay Octopuce: Total Cluster Charge
Some differences in amplification gaps between different chips are seen by microscope
(studies are not conclusive, might come from different thickness of the glue under the chips)
Chips 1 and 4 have a lower response (same trend as for the cluster size distribution) 
Difference in amplification gap
Difference in threshold (too big to explain differences between 1 &4 and others)
After equalization:
A.Chaus, D. Attie

## Slide 16
DESY Test-Beam with Saclay Octopuce & Large Prototype TPC (Dec. 2010)
B = 0 T
He/iC4H10(80:20)
B = 1T
He/iC4H10(80:20)
4 chips
2 chips

## Slide 17
Timepix pane for Large Prototype TPC
Mezzanine board
Mother board
Guard ring
Chips on a mezzanine board making wire bonding easier
Large Prototype compatible
Heat dissipator
Saclay Octopuce in the Large Prototype TPC

## Slide 18
Saclay Octopuce in the Large Prototype TPC at 0 T
He/Iso 80/20
Vmesh = 380V
Time mode
Shutter Time: 100 ms, start given by beam trigger
Occupancy is very low:
< 0.1 % (no noisy frames)
Example of a typical event:

## Slide 19
ROOT based analysis package developed by John Idarraga (LAL)
C++ classes including processors:
  OctoCEA (define in a few minutes)
  Pattern recognition of tracks for low threshold
MAFalda: Medipix Analysis Framework
Use MAFalda as the first step  implement reconstruction algorithm in Marlin TPC

## Slide 20
MAFalda: Cluster Finding and Track Segments (I)
Form cluster from the pixels with “discontinuity” < 40 pixels (each cluster
should contain > 12 pixels)
Calculate the linear regression with all points in the cluster (red dotted line)
Calculate residuals from the red line to each point (if > 80 % the points are within 20 pixels – this is the “track segment” !
A. Chaus

## Slide 21
MAFalda: Cluster Finding and Track Segments (II)
Check incident angles of track segments
 Take segments within +/-10 degrees
(electron beam is in the horizontal direction)
Number of segments
per reconstructed track
(most tracks consists of
2-3 segments):
A.Chaus

## Slide 22
Track Reconstruction
Y = ax +b
b : 417.558 | a : -0.021984
b : 421.54 | a : -0.0199386
b : 428.789 | a : -0.0341787
b : 431.063 | a : -0.0362019
Reconstruct track from track segments (if more than 2 segments):
Apply data quality cuts to make sure all segments correspond to a given track
Perform a linear regression for all pixels on a track  reconstructed track (blue dotted line)
A.Chaus
A. Chaus

## Slide 23
Track Reconstruction: Residuals
An unbiased estimate of the single point resolution: s = Sqrt(s1 * s2):
Perform fit to all pixels and calculate the distance between each pixel and the position
of the point of closest approach along the fitted track:
s1 = Gaussian fit, when all pixels are included in the track fit
s2 = Gaussian fit, omitting pixel under consideration from the track fit
in number
of pixels
A. Chaus

## Slide 24
DESY Test-Beam: Track Residuals
Summary of residuals (s = sqrt(s1*s2)) for different Z-coordinates:
(corresponds to the different positions of electron beam
passing through the Large Prototype TPC)
A. Chaus

## Slide 25
DESY Test-Beam: Saclay Octopuce
Track Residuals as a function of Z
(Large Prototype TPC)
PRELIMINARY:
PRELIMINARY:
Very low statistics, behavior needs to be understood
A. Chaus

## Slide 26
History: “Octopuce”s with Timepix Readout
2010: Saclay “Octopuce” (InGrid/MM)
Module  studied in lab at DESY test-beam
2013: NIKHEF “Octopuce”(InGrid/MM):
some problems with shorts & readout
2013: Bonn 2 “Octopuce” Modules
(Ingrid/MM) and 3-GEM  studied with SRS Readout at DESY Test-Beam in April 2013
Saclay
Saclay
Bonn
X-Rays

## Slide 27
Two Bonn “Octopuce”s (Triple-GEM & MM / InGrid) Studies at DESY Test-Beam
4 Timepix ASICs
MM/InGrid Track Reconstruction: Residuals
Studies in LCTPC Prorotype:
Different magnetic fields
(0 and 1 T)
Different drift fields
4 Timepix ASICs
Triple-GEM
MM/InGrid

## Slide 28
Bonn “Octopuce” (MM / InGrid) Studies at DESY Test-Beam
B = 1 T; E = 230 V/cm
Theoretical limit:
σ = Dt√z,
Dt = 106 (Magboltz)
B = 0 T; E = 230 V/cm
Theoretical limit:
σ = Dt√z,
Dt = 320 (Magboltz)
From fit (exp.):
0157.3
Dt = 100.3

From fit (exp.):
0~ 0
Dt = 304.6

Analysis in
MAFALDA
FRAMEWORK:

## Slide 29
Pixel Readout for LC TPC: Present and Future
Short-Term Plans:
Improve SU8-based « InGrid « structure (less dead area, higher yield, protection)
Better understand and improve field distortion issues at the « InGrid » plane
(in between chips and at the borders of the Octopuce);
Develop simulation chain to compare momentum resolution, double track resolution, dE/dx and pattern recognition to pad-based readout and to optimize the geometry (e.g. minimize the needed chip coverage; improve electric field between chips and at the chip borders)
Long-Term Plans:
Improved chip functionality  Timepix3  transfer InGrid process to Timepix3
Adapted readout system  high-speed SRS readout for Timepix(3)(8 &100 chips);
Improved mass production of InGrids  in collaboration with IZM, Berlin
Work on more realistic cooling and power pulsing;
Develop a backside connection technology both for LV/data and for the HV;