## Slide 1: COsmic Ray Telescope at Orsay (CORTO)
Andrii Natochii
CORTO team:
V. Puill, L. Burmistrov, V. Chaumat, D. Breton, F. Campos, C. Cheikali, J. Maalmi, J. Peyre, A. Stocchi, C. Sylvia, J.F. Vagnucci, E. Bals, B. Genolini, X. Grave, Mimae Kim, D.W. Kim
Taras Shevchenko National University of Kyiv
The Linear Accelerator Laboratory at Orsay

## Slide 2: Main steps of the presentation:
Introduction
CORTO
  MRPC
Last results
Problems
Upgrade
New results
Efficiency map of the MRPC
2D Scanning Device for CORTO
  First prototype
  Final version
  Software and controlling system
Conclusion

## Slide 3: Introduction
How to test a new detector in HEP at first steps and to characterize its response function?
Radioactive Sources
Particle Accelerator
Nuclear Reactor
Laser
Cosmic Rays

## Slide 4: CORTO
CORTO (2013)
Structure:
1) two positional sensitive strip gas detectors (MRPC)
(1 m x 2 m x 5 cm), which provide the track reconstruction.
2) read-out electronics: NINO Application-Specific Integrated Circuits (ASIC), WaveCatchers, interface cards.
3) gas supply system for MRPCs.
Main task is to reconstruct muon tracks.

## Slide 5: MRPC (Multigap Resistive Plate Chamber)
Provide the track reconstruction

## Slide 6: Last results (2013)
A time resolution of 250 ps (for a bias voltage of 16 kV), coordinate resolution along the X direction of 1.8 cm and along the Y direction of 2.1 cm.
Last configuration:
- Two MRPCs.
- One read-out board of WaveCatcher
(48 channels).
- Signals rate is ~ 100 Hz for coincidences between those two MRPCs.
The two-dimensional histograms of the reconstructed position of “Short” (top picture) and “Long” (bottom picture) detectors.

## Slide 7: Problems
Too high rate of MRPC signals (should be ~ 50 Hz).
Not nominal time resolution (should be ~ 100 ps).
Not uniformly distribution of efficiency on each MRPC:
Reconstructed hits from bottom MRPC (2013).
SOLVED
SOLVED
IN PROCESS

## Slide 8: Upgrade
In 2014 was added third MRPC for CORTO.
Three boards of WaveCatcher (total number of the channels is 48 x 3 = 144).

## Slide 9: New results
CORTO scheme (2014)
Reconstruction of Middle MRPC position using Top and Bottom MRPCs (without channels calibration).

## Slide 10: Efficiency map of the MRPC
Still have few gaps of the MRPC efficiency =(
Reconstructed hits from bottom MRPC (2014).
Solution
Efficiency checking

## Slide 11: 2D Scanning Device for CORTO (2D SDC)

## Slide 12: First Ukrainian Prototype

## Slide 13: Final version of 2D SDC at LAL
Main parameters:
- Position accuracy less than 0.1 mm
- Linear speed around 5 mm/s;
- Maximum tested weight is 20 kg

## Slide 14: Software and controlling system
Software window (GUI).
Arduino UNO with Ethernet Shield
Controlling components:
Two stepper motors (NEMA23), two microstepping drivers (LeadShine), two power supplies (5V/3A & 24 V/5A), microcontroller (Arduino Uno), Ethernet Shield for ethernet connection and remotely controlling.

## Slide 15: Conclusion
During this year we have done a lot of important steps for CORTO improvement:
Third MRPC, Three WaveCatcher cards, 2D SDC
All this things will help us to prepare the CORTO for it’s main mission – to test different types of detectors, knowing MRPC accuracy and efficiency.

## Slide 16: Acknowledgment
Thanks a lot for helping and professional advice:
Véronique Puill
Leonid Burmistrov
Jean François Vagnucci
Dominique Breton
All CORTO team
And guys from central LAL engineering atelier and from LAL tool shop