## Slide 1
Detection properties and internal activity of newly developed La-containing scintillator crystals
G. Hull1, F. Camera2, G. Colombi2, M. Josselin1
Institut de Physique Nucléaire d'Orsay, CNRS-IN2P3, Univ. Paris Sud, Univ. Paris-Saclay, Orsay France
2 Università degli Studi di Milano, INFN Sezione di Milano, Milan, Italy
LIA France-Ukraine Workshop 2016 September 27th, 2018

## Slide 2
La-containing crystals
The tested crystals and the experimental set-up
Light yield and energy resolution
Response vs interaction point of the γ radiation
Internal activity
Summary
Outline
Pag. 1

## Slide 3
Pag. 2
New scintillator crystals
Fast (~16 ns)
Bright (>60kγ/MeV)
LaBr3:Ce
Fast (~17 ns)
Bright (>60kγ/MeV)
No internal activity
CeBr3
SrI2:Eu
Very bright (>100kγ/MeV)
Slow (~1μs)
Suffer of self-absorption
CLYC
(Cs2LiYCl6:Ce)
Good energy resolution
Sensitive to thermal n
n/γ PSD discrimination

## Slide 4
Pag. 3
New La-based scintillator crystals
Co-doped LaBr3:Ce and Cs2LiLaBr6:Ce
 | Light Yield (ph/MeV) | Decay Constant (ns) | Wavelength of emission (nm) | ΔE/E @ 662 keV (%) | Propor.
LaBr3:Ce,Sr | >80,000 | 25 | 390 | ~2.6 | Good
CLLB Cs2LiLaBr6:Ce | 60,000 | 180, 1080 | 410 | ~3.0 | Good
JR1-PASPAG: Phoswich scintillator assemblies: Application to the Simultaneous detection of Particle and Gamma radiation
TASK1 Novel Scintillator Materials
CLLB and LaBr3:Ce,Sr crystals available from Saint Gobain

## Slide 5
Pag. 4
The experimental set-up
CLLB (CLLB1)  1  1
CLLB(CLLB2)  2  2
LaBr3:Ce,Sr  1.5  1.5
LaBr3:Ce  1  1
PMT R6231-100-SEL-MOD
 2
8 dynodes
SBA photocathode
G = 4.7E4 @ 750V
G = 7.2E4 @ 800V
Standard spectroscopic chain:
  Mesytech preamplifier
  Ortec spectroscopy amplifier
  Multichannel analyser (MAESTRO and MCA8000A)
Digitizer CAEN DT5730
  8-channel 14-bit 500MS/s

## Slide 6
Pag. 5
Decay time
Collected charge as a function of the integration time (digitizer’s gate width)
y=A0 + A1[1-e(-x/τ1)] + A2[1-e(-x/τ2)]
CLLB
LaBr3:Ce,Sr
τ = 30.80 ± 0.01 ns
τ1 = 154.39 ± 0.04 ns 63%
τ2 = 1096.8 ± 0.3 ns 37%

## Slide 7
Pag. 6
Light yield and linearity
LY measured with the digitizer in the energy range between 60 keV and 1.77 MeV
Light yield proportion
LaBr3:Ce : LaBr3:Ce,Sr : CLLB1 : CLLB2
1 : 1.29 : 0.78 : 0.63
Chart: Energy [keV]
| 662.0 | 511.0 | 1275.0 | 1173.0 | 1332.0 | 31.0 | 81.0 | 276.0 | 303.0 | 356.0 | 122.0 | 244.0 | 344.0 | 411.0 | 444.0 | 779.0 | 867.0 | 964.0 | 1408.0
CLLB1 | 5625 | 4340.4861 | 10793.75 | 9990.2778 | 11338.8889 | 278.4722 | 697.5347 | 2352.9514 | 2569.7917 | 3022.9167 | 357.5486 | 1031.5069 | 2056.875 | 2894.6736 | 3450.9722 | 3910.1389 | 6529.8611 | 7242.2917 | 8076.25 | 11789.5139
CLLB2 | 4458.6806 | 3451.25 | 8571.5278 | 7945.1389 | 9035.4167 | 542.5347 | 2024.8264 | 2394.7917 | 818.8611 | 1648.0556 | 2326.9444 | 2775 | 2982.6389 | 5265.2778 | 5818.0556 | 6511.1111 | 9523.6111
LaBr3:Ce | 7361.1702 | 12807.1809 | 14462.766 | 5675.1064 | 13853.9894 | 6330.8511 | 11661.4362 | 18929.5213 | 3081.6489 | 3375.266 | 3967.8191 | 4276.5957 | 1351.8351 | 2732.5798 | 3841.383 | 4591.4894 | 4951.3298 | 8607.7128 | 9569.4149 | 10609.8404 | 15261.7021 | 813.0585 | 936.7287 | 889.4947
LaBr3:Ce,Sr | 9292.0213 | 16455.8511 | 18632.9787 | 7175.2926 | 17700.5319 | 8048.75 | 14923.9362 | 24500.7979 | 3888.2979 | 4261.7021 | 5006.0904 | 5402.9255 | 1716.4628 | 3443.2181 | 4838.2447 | 5777.1277 | 6236.7021 | 10888.5638 | 12107.9787 | 13451.8617 | 19486.4362 | 1052.0745 | 1210.4255 | 1143.0053
Deviation from linearity less than 1% for all tested crystals

## Slide 8
Chart: Energy [keV]
| 276.0 | 302.0 | 356.0 | 511.0 | 662.0 | 1173.0 | 1275.0 | 1332.0
LaBr3:Ce 1"x1" | 4.4581 | 3.974 | 4.1234 | 3.7836 | 3.1701 | 2.4399 | 2.4498 | 2.1384
LaBr3:Ce,Sr 1.5"x1.5" | 3.0581 | 3.3805 | 3.1885 | 2.9357 | 2.3901 | 1.8077 | 1.753 | 1.7277
CLLB 1"x1" | 8.0914 | 5.5172 | 5.2438 | 4.3758 | 4.1561 | 3.2453 | 2.9413 | 2.6338
CLLB 2"x2" | 7.2457 | 6.0658 | 6.0519 | 5.6621 | 5.383 | 4.0875 | 4.3063 | 4.2365
Pag. 7
The energy resolution
R measured with a standard spectroscopic chain in the energy range between 276 keV and 1.3 MeV
R=4.2% CLLB1
R=5.4% CLLB2

## Slide 9
Pag. 8
The energy resolution
CLLB 2”x2” – FWHM = 35.7 keV
CLLB 1”x1” – FWHM = 27.5 keV
LaBr3:Ce – FWHM = 21 keV
LaBr3:Ce,Sr – FWHM = 15.9 keV

## Slide 10
Response vs irradiation point
Pag. 9
Highly-collimated 400MBq 137Cs source mounted on a support that allows rigid translations in two directions
  Collimator in heavymet 8 cm long –  1 mm large
  A 60Co source used as reference
Scan along the X and Y axis:
6 4-mm step for the CLLB1 and 12 4-mm step for the CLLB2
PMT
Scan along the z axis: 6 4-mm step for the CLLB1 and 13 4-mm step for the CLLB2

## Slide 11
Response vs irradiation point
Pag. 10
Scan along the X and Y axis
Less than 1 ch variation
Chart: Source postion [mm]
| -23.0 | -19.0 | -15.0 | -11.0 | -7.0 | -3.0 | 1.0 | 5.0 | 9.0 | 13.0 | 17.0 | 21.0
Cs137 Peak along Y | 228.51 | 227.65 | 228.47 | 229.19 | 229.75 | 229.72 | 229.81 | 229.5 | 229.04 | 228.88 | 228.91 | 228.66
Cs137 Peak along X | 225.56 | 226.3 | 226.8 | 227.64 | 229.64 | 228.03 | 227.88 | 227.9 | 228.1 | 228.4 | 228.3 | 227.47
CLLB 2”x2”
Chart: Source postion [mm]
| -10.0 | -6.0 | -2.0 | 2.0 | 6.0 | 10.0
Cs137 Peak along Y | 173.67 | 173.62 | 173.59 | 173.37 | 173.31 | 173.77 | 173.555
Cs137 Peak along X | 173.92 | 173.72 | 173.74 | 173.6 | 173.94 | 174
CLLB 1”x1”
Between 3 and 5 ch variation, mostly on the edges

## Slide 12
Chart: Source position along Z [mm]
| 4.0 | 8.0 | 12.0 | 16.0 | 20.0 | 24.0
CLLB 1"x1" | 168.9 | 166.9 | 165.79 | 165.11 | 164.43 | 164.54
Chart: Source position along Z [mm]
| 4.0 | 8.0 | 12.0 | 16.0 | 20.0 | 24.0 | 28.0 | 32.0 | 36.0 | 40.0 | 44.0 | 48.0 | 52.0
CLLB 2"x2" | 231.1 | 228.84 | 227.46 | 226.24 | 225.17 | 224 | 223.23 | 222.01 | 221.31 | 220.59 | 220.15 | 219.97 | 220.02
Response vs irradiation point
Pag. 11
CLLB 1”x1”
CLLB 2”x2”
5 ch
11 ch
Scan along the Z axis
Evident anisotropy in the CLLB response
Contribution to the energy resolution:
CLLB1 = 17 keV ( R=2.6%)
CLLB2 = 29 keV ( R=4.4%)
Subtracting this contribution to the measured energy resolution:
R = 3.3% for CLLB1 and R = 3.1% for CLLB2

## Slide 13
Peak width vs irradiation point
Pag. 12
CLLB1: collimated source FWHM < 23 keV
R ~ 3.3%
Chart: Irradiation point distance from the PMT [mm]
| 4.0 | 8.0 | 12.0 | 16.0 | 20.0 | 24.0 | 28.0 | 32.0 | 36.0 | 40.0 | 44.0 | 48.0 | 50.0
FWHM - CLLB2 | 46.5 | 36.9 | 32.9 | 29.9 | 30 | 28.6 | 29.9 | 28.4 | 28.3 | 28.1 | 27.3 | 28.2 | 26.2
FWHM - CLLB1 | 22.324 | 20.613 | 22.4055 | 23.0573 | 22.7314 | 21.9166

## Slide 14
Pag. 13
138La activity
138La: γ (1436 keV + 789 keV) + β
227Ac: mainly α (5.0 – 7.4 MeV)

## Slide 15
The internal activity
Pag. 14
227Ac α peaks in the different regions for the different crystals
Higher α-activity for the co-doped LaBr3:Ce compared to the standard one
Higher α-activity for the CLLB1 with respect to the CLLB2
Peaks at 1470 and 1440 keV well separated for the LaBr3:Ce,Sr
LaBr3:Ce
LaBr3:Ce,Sr
CLLB1
CLLB2
~1470 keV
=
1436 keV
+
K-α Ba
~1440 keV
=
1436 keV
+
L and M Ba

## Slide 16
The internal activity
Pag. 15
Crystal | Internal countig rate [Cts/s/cm3] E>70keV | α contribution [%] | α energy region [MeV GEE]
LaBr3:Ce | 0.53 | 1.1 | 1.7 – 2.7
LaBr3:Ce,Sr | 0.84 | 11.9 | 2.2 – 3.2
CLLB1 | 0.31 | 11.8 | 2.8 – 4.7
CLLB2 | 0.32 | 6.4 | 2.8 – 4.7
LaBr3:Ce,Sr has more self-activity than LaBr3:Ce
Different quenching for the different scintillators
Different alpha contribution

## Slide 17
Pag. 16
Summary
We tested two CLLB crystals, with different volume, and a co-doped LaBr3:Ce from Saint Gobain
LaBr3:Ce,Se is brighter and shows better energy resolution than LaBr3:Ce
The CLLB1 produces 25% more light than the CLLB2
We observed a strong anisotropy in the CLLB light yield, which affects the measured energy resolution
The alpha quenching is different for all the tested crystals, in particular it is higher for LaBr3:Ce and so the 227Ac contribution is better separated for the co-doped sample
The alpha contribution is higher for the CLLB1 than for CLLB2, this might indicates an improvement of the materials purification in the crystals production process
We are carrying on an R&D activity on new scintillator materials for nuclear physics applications in the framework of the ENSAR2-JR1-PASPAG project

## Slide 18
Acknowledgment
This work was supported by the H2020-INFRAIA-2014–2015 Grant Agreement 654002 – ENSAR2-PASPAG

## Slide 20
Pag. 15
AmBe(Ni) n/γ PSD distrimination
PSD =
Ql-Qs
Ql
Charge comparison method:
Ql calculated over a long gate = 4 μs,
Qs calculated over a short gate = 200 ns
CLLB 2”x2”
CLLB 1”x1”

## Slide 21
Pag. 16
AmBe(Ni) n/γ PSD distrimination
FoM =
Peakn- Peakγ
FWHMγ+FWHMn
CLLB 2”x2”
Separation from n to γ-rays, but not between n and α
γ
n/α
CLLB 1”x1”
γ
n/α
FoM=1
FoM=1.7

## Slide 22
Peak position vs irradiation point
Pag. 12
0 mm
8 mm
16 mm
24 mm
32 mm
40 mm
48 mm
CLLB 2”x2”
Chart: Source position along Z [mm]
| 0.0 | 4.0 | 8.0 | 12.0 | 16.0 | 20.0 | 24.0 | 28.0 | 32.0 | 36.0 | 40.0 | 44.0 | 48.0 | 52.0
FWHM | 46.5 | 36.9 | 32.9 | 29.9 | 30 | 28.6 | 29.9 | 28.4 | 28.3 | 28.1 | 27.3 | 28.2 | 26.2
R | 7 | 5.6 | 5 | 4.5 | 4.5 | 4.3 | 4.5 | 4.3 | 4.3 | 4.2 | 4.1 | 4.3 | 4
Energy resolution -R [%]
Collimated source
R ~ 4.3%
FWHM < 30 keV
Collimated
Uncollimated

## Slide 23
Pag. 5
Decay time
Collected charge as a function of the integration time (digitizer’s gate width)
y=A0 + A1[1-e(-x/τ1)] + A2[1-e(-x/τ2)]
Gate > 4 μs
CLLB
LaBr3:Ce,Sr
Gate > 500ns

## Slide 24
Pag. 15
n peak detected at ~3.3 MeV GEE for the two CLLB crystals
Good linearity up to >4.4 MeV for all the testes crystals
Peak @ 9 MeV well visible for the CLLB2
Saturation above 7 MeV for the LaBr3:Ce
Irradiation with neutrons and g-rays up to 9 MeV
Response to the AmBe(Ni) source
n

## Slide 25
Pag. 16
n/γ PSD distrimination – AmBe
PSD =
Ql-Qs
Ql
Charge comparison method:
Ql calculated over a long gate = 4 μs,
Qs calculated over a short gate = 200 ns
CLLB 2”x2”
CLLB 1”x1”

## Slide 26
Pag. 17
n/γ PSD distrimination - AmBe
FoM =
Peakn- Peakγ
FWHMγ+FWHMn
CLLB 2”x2”
Separation from n to γ-rays
γ
n
CLLB 1”x1”
γ
n
FoM=1
FoM=1.7